## Slide 1
Smart Cognitive-Motor Rehabilitation Mat for Remote Exercise Monitoring
Electrical & Computer Engineering
Group 47: Scott Lopez, Jashan Virdi, Adithya Balaji
5/4/2025

## Slide 2
Introduction
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 3
IMAGE / GRAPHIC
What is Multiple Sclerosis (MS)?
MS affects the protection around nerves
Loss of protection creates scar tissue
Hinders balance, mobility, and strength
How can it be treated?
Physical exercise has shown improvements in balance, mobility, and strength; however, most forms of physical exercise are not accessible to older adults with MS.
Introduction
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Scott

## Slide 4
IMAGE / GRAPHIC
Solution
The Smart Cognitive-Motor Rehabilitation Mat for Remote Exercise Monitoring (Smart Mat) solves this problem by making square stepping exercise accessible and safe at home for older patients.
Introduction
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Scott

## Slide 5
IMAGE / GRAPHIC
First Prototype
The first prototype allows for the mat to be carried and foldable. It used Velostat and copper foil to detect pressure.
Problems
Sensors are unreliable, often provide incorrect feedback
Requires Wi-Fi in order to function
Uses a 12 V adapter due to LED voltage requirement
Using devkit, breadboard
Introduction
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Jashan

## Slide 6
IMAGE / GRAPHIC
Second Prototype
We aimed to make our new prototype more reliable, have Bluetooth capability, and power it with USB-C.
High Level Requirements
Reliability: Improve reliability of sensors to withstand a minimum of thirty rolls.
Data Transmission Latency: The system must provide real-time feedback with a data transmission delay of no more than 500 ms between the smart mat and the mobile device.
Power Source: Run the mat on a 5 V USB-C to allow ease of replacing the power adapter.
Introduction
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Jashan

## Slide 7
Design
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 8
IMAGE (Send to back so Block I is not covered.)
Block diagram with three main subsystems: Mat, Power, and Data Processing.
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Adithya

## Slide 9
IMAGE / GRAPHIC
Velostat Designs
The top left image is the previous design. We initially wanted to switch to the design on the bottom left, a Velostat sandwich.
Issues
We tested several designs with velostat, but none had consistent results. After thorough testing we concluded that Velostat wasn’t reliable, and mechanical switches were needed.
Design
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Adithya
This design would increase the area of a Velostat between both conductors, and so it would better sense when pressure was applied to that square.

## Slide 10
IMAGE / GRAPHIC
Mat Subsystem
Non-slip yoga mat and cloth for wire protection
4 by 8 (x,y) grid of buttons
5 columns of LEDs
R&V requirements
Can roll at least 30 times and sensors aren't affected
The incorrect squares should not be activated if not stepped on
Velostat should accurately detect if a person has stepped on them
Design
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Scott

## Slide 11
IMAGE / GRAPHIC
Data Processing Subsystem
Row and column scanning using shift register and multiplexer (MUX)
ESP32-S3 does grid calculation
R&V requirements
Should correctly load shift register to keep correct scanning frequency
Should send data during routine within 500 ms and accurate
Connected device should be able to transmit data across a 10 meter room
Design
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Scott

## Slide 12
IMAGE / GRAPHIC
Power Subsystem
USB-C receptable and linear regulator
5 V for LEDs, 3.3 V for ESP32-S3 and ICs
R&V requirements
System should get 5 ± 0.3 V from the USB-C port, and be able to draw enough current for the microcontroller and the LEDs (nominally 2.66 A)
System should be able to provide 3.3 V ± 0.2 V to the ESP32-S3
Design
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Jashan

## Slide 13
IMAGE / GRAPHIC
PCB Layout
87 mm by 90 mm (not including USB-C receptable and antenna)
Notes
Screw terminals to securely connect LEDs
Headers to connect to row and column wires
USB programming for ESP32-S3, also included the UART programming circuit for alternate programing
Design
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Jashan

## Slide 14
PCB Schematic
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Jashan

## Slide 15
IMAGE / GRAPHIC
ESP32-S3-Wroom-1 Program
ESP32-S3 was used to perform the following tasks
Design
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
BLE connection with external design
Pre-program routines for workouts
Process row and column scanning logic
Programming flowchart
Acts as a BLE GATT server, advertising ESP32-S3
StepMatDefines a custom service + characteristic for 2-way messaging
Responds with BLE notifications for real-time updates (latency <60ms)
Stores any number of routines as arrays of sensor steps
Automatically tracks correct vs incorrect steps
Logs time, accuracy, and total performance
Uses 74HC595 shift register to activate columns (COL0–COL3)
Reads row sensor state using CD74HC4067 multiplexer
Scans all 32 sensor positions every loop cycle (~5ms delay)
Notes: Adithya

## Slide 16
App
Another visual representation of the mat, written in HTML. The app also allows for easy selection of routines.
Design
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Visual depiction of sensor to press highlighted in blue and real time updates on HTML page
Minimalist landing page before starting workout
Home page to select workout from drop down list and connect to BLE before starting
Information on workout results given for user for more feedback
Notes: Adithya

## Slide 17
Building and
Testing
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 18
IMAGE / GRAPHIC
Breadboard and Mock Demo
The breadboard demo used a 2 by 2 velostat grid with one LED strip down the center.
The Mock demo was a 2 by 2 button grid with the PCB to show PCB functionality.
Differences
Velostat to button
PCB instead of devkit
Building and Testing
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Scott

## Slide 19
Velostat Prototype
A 2 by 6 grid on a yoga mat using Velostat.
Features
Screw terminals to easily connect LEDs
Headers to allow for disconnection of row and column for PCB testing
UART for alternate programing
Problems
Velostat was inconsistent
Few squares
Building and Testing
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Scott

## Slide 20
Velostat Designs
Three designs of Velostat were used to try and reduce inconsistency.
Types
Copper foil on Velostat
Copper foil folded on wire
Aluminum foil
Building and Testing
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Scott

## Slide 21
Codebase testing
Serial Monitor Debugging: Used and BLE logs to verify sensor reads
Unit Testing Components Separately: Mux, Shift register
Manual Sensor Validation: Verified each sensor position individually by stepping on them
BLE Roundtrip Latency Testing: Implemented mechanism to actively measure Bluetooth latency between ESP32-S3 and browser
Routine Execution Tests: Ran through each of the 10 routines and verified handling
Browser Integration Debugging: Tested the HTML dashboard for accurate display
Edge Case Handling: Simulated partial steps, back-to-back presses
Building and Testing
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Adithya

## Slide 22
IMAGE / GRAPHIC
Power Testing
Testing 5 V and 3.3 V outputs with multimeter
Testing current draw by using DC power supply to power board
Building and Testing
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Jashan

## Slide 23
Final Demo
A 4 by 8 grid on a yoga mat using buttons.
Features
Similar to Velostat test
Table cloth to cover wires
Duct tape to show where buttons are located
Building and Testing
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Jashan

## Slide 24
CHART / GRAPH
Video of project
Building and Testing
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 25
Ethics
Data privacy and security
Responsible development and testing
Safety
Electrical Safety
Reliability and accuracy
Ethics and Safety
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Adithya

## Slide 26
IMAGE / GRAPHIC
What we learned:
Mechanical devices are more reliable for certain tasks
Debugging PCBs
Timeline of orders can affect project
What we would do differently:
Use different headers and/or connectors for the row and column wires
Scrap velostat from the start
Future Work
Implement design on full mat
Replace buttons with other sensors if desired
Change connectors to clean up wires
Add real square stepping routines to program
Test with real end users
Conclusion
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING
Notes: Adithya

## Slide 27
Thank You
Group 47: Scott Lopez, Jashan Virdi, Adithya Balaji
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING