## Slide 1
EEG Drowsiness Detection Device
Electrical & Computer Engineering
By (Team 28): Nikhil Talwalkar and Senturran Elangovan
December 9th 2025

## Slide 2
Problem
Drowsiness is a major safety risk in jobs involving sustained attention, such as night-shift work. NHTSA data links tens of thousands of deaths to drowsy driving.
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 3
Solution
A wearable EEG-based drowsiness detector that continuously monitors brain activity to determine drowsiness. Capture real-time EEG data, filters and processes using ADS1299/STM32, and provides buzzer alerts in the event of sleep detection.
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 4
EEG acquisition: Device must capture EEG signals with sufficient accuracy, tested through eye blinks
Real-time drowsiness detection: The system must detect micro-sleeps or drowsiness with at least 90% accuracy and a low false positive rate
Prompt alerting: The alert mechanism (buzzer) must activate within a timely manner) from the onset of detected drowsiness
Wearability and safety: The device must remain safe and comfortable for at least 8 continuous hours. We will verify this through randomly selected volunteers and quantify comfort through a number scale (1-5)
High Level Requirements
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 5
High Level Block Diagram + Visual Aid
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 6
Schematic subsystems
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 7
PCB Progress
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 8
PCB version 1. Our TL431 design was made with the wrong footprint, therefore changes were made for version 2.
PCB Progress
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 9
PCB version 2. The BMS_Out (GND) pin was connected to another net when updating the schematic to PCB.
PCB Progress
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 10
PCB version 3. Inductor rating mismatch, so new power inductors required a larger footprint.
PCB Progress
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 11
Power Subsystem
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 12
BMS requirements table and subsystem schematics.
Requirements and Schematics
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 13
The BMS maintains a voltage less than 4.2V
We tested charging the battery using 8V power supply and 200mA. 4.185V was seen supplied to the battery (in line with our requirements)
Our Buck/Boost converter regulates 3.7V at 3.3V and 5V
3.89V battery voltage measured
Buck converter output: 3.048V – 7.6% error
Boost converter output: 5.077V – 1.6% error
BMS & Converters
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 14
Our target battery life was 8 hours.
To estimate battery life we assume linearity across operational voltage 4.2V – 3.3V. On a 10-minute average, boost converter drained 0.02V.
Average current draw of our system is 400mA. Therefore, a fully charged 3000mAh battery should last 7.5 hours – 6.25% below 8 hours.
Battery Life
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 15
Analog + Control Unit Subsystems
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 16
IMAGE / GRAPHIC
ADS act as the Analog Front End, while the STM does DSP and decision logic
ADS1299:
24-bit ADC with 250Hz sampling
3 electrodes configured over SPI1
STMF412RGT6:
ARM Cortex-M4 Microcontroller
Digital processing for real-time monitoring and sleep detection
STM/ADS1299
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 17
STM/ADS1299 Requirements
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 18
Collects data at 250 samples per second
Bias buffer for lower noise
CH4 with gain 24x for amplification
Other channels all powered down to reduce noise and current
Reads full 24bit frame and converts to volts
Sampling and Digital conversion
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 19
Sampling frequency of all 13 test subjects. Plotted with the total number of samples in each run.
Sampling and Digital conversion
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 20
Use 256-sample windows with 50% overlap
Remove DC offset (mean normalization
Total power in the window
Power for theta (6Hz) and alpha (10Hz) using Goertzel’s algorithm
Alpha (8-13Hz): Active when a person is relaxed with eyes closed
Theta (4-7Hz): Increases as person goes from drowsiness to stage-1 sleep
Power Ratios
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 21
Power Ratios
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 22
Window ratio of all 13 test subjects. Plotted over the no of seconds (windows).
Power Ratios
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 23
Baseline calibration of open eyes during the first 10-15 seconds of startup
Average ratio of the first ten window samples
Helps ignore high ratios (>4)
Baseline ratio is logged on debugger
Dynamic threshold uses baseline and tuned value of 1.8
Baseline and Dynamic Thresholds
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 24
Ratio plot of one test subject during trials (Anya).
Baseline and Dynamic Thresholds
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 25
Ratio plot of one test subject during trials (Arushi).
Baseline and Dynamic Thresholds
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 26
Action Subsystem
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 27
Flowchart of decision algorithm.
Decision Logic and Buzzer Alerts
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 28
This is the main state machine logic in the main loop
The loop maintains a drowsy_counter for each window
If drowsy, the counter is incremented and decremented if not
Alerts is reached when:
  drowsy_counter >= 2
When alert is triggered, buzzer on for 500ms
Event Log is created .
Decision Logic and Buzzer Alerts
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 29
Response time of STM processing of all 13 test subjects.
Decision Logic and Buzzer Alerts
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 30
We asked all test subjects a series of questions after the trial
Evaluated wearability with a 1-5 scale across 13 subjects
6 out of 13 participants rated comfort as a 4 or 5
Only 3 reported a high pressure of 4
Wearability
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 31
Regulated 3.7 nominal voltage to 3.3V and 5V with less than 10% error
BMS battery protection capped voltage at 4.2V (600mA)
Reliable EEG streaming: Constant sampling rate of 250Hz, very few data overruns, and very fast response time (3-8ms)
Drowsiness detection behaves as predicted: No constant buzzing, clear evidence of alerts above threshold ratio
Wearability close to target and buzzer is distracting
Dips in sampling frequency
Frequency response not fully characterized
8-hour battery life not fully verified (theoretical)
UART on stm was disabled
Successes, Challenges and Verification
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING

## Slide 32
Learning Outcomes
What to do differently
Recommendation for future work
Conclusion
GRAINGER ENGINEERING
ELECTRICAL & COMPUTER ENGINEERING