## Slide 1
Team58: AdheraScent
Megan Shapland, Wenchang Qi

## Slide 2
Introduction to your team and your project.

## Slide 3
AdheraScent: A smart pill dispenser using scent-based notifications to improve medication adherence.
Core Hardware: ESP32-S3 microcontroller manages schedules and triggers an ultrasonic aroma diffuser.
Local Connectivity: Hosts a dedicated web server
(SoftAP mode) for easy scheduling via smartphone without internet.
Project Introduction

## Slide 4
Objective. What problem are you solving?

## Slide 5
Age-Related Decline: Assists users with diminished vision, hearing, and cognitive functions in managing complex regimens.
Notification Fatigue: Bypasses traditional digital "noise" by using a distinct, non-intrusive olfactory reminder.
High Technical Barriers: Eliminates complex setups with automated time synchronization and an intuitive web interface.
Problems Solved

## Slide 6
Brief review of original design, statement on areas of design that changed, and overview of each functional block's requirements.

## Slide 10
Design Changes

## Slide 12
Power Subsystem

## Slide 13
Requirements
Battery Charing Current 500 mA ± 50 mA
Output 3.3V ± 0.15V
Output 6.0V ± 0.15V
Cut off power output if the battery voltage drops below 2.5V
Power Subsystem
Verification Result
PASSED- Measured by Digital Multimeter
PASSED- Measured by Digital Multimeter
PASSED- Measured by Digital Multimeter
PASSED- Tested by letting battery voltage drops below 2.5V

## Slide 14
Control Subsystem

## Slide 15
Control Subsystem
Requirements
Microcontroller must trigger interrupts within 200ms of Hall sensor changes
Output 3.3V ± 0.15V
ESP32 must communicate with RTC at 100 kHz ± 5 kHz frequency
Sync system clock via web client within one-minute post-connection
ADC must calculate battery voltage within ± 0.1 V of true value
Requirements
PASSED- Measured by Observation
PASSED- Measured by Digital Multimeter
PASSED- Measured by Oscilloscope
PASSED- Measured by Observation
PASSED- Measured by Digital Multimeter

## Slide 16
Actuation & Sensing Subsystem

## Slide 17
Actuation & Sensing Subsystem
Requirements
MOSFET must fully saturate V_DS ≤ 0.15 V under 3.3V logic signal
Pull-down resistor R5 must keep gate voltage ≤ 0.1 V when floating
Hall sensors must achieve 100% detection accuracy over 50 trials
LC network must suppress power ripples from the atomizer’s draw
Requirements
PASSED- Measured by Digital Multimeter
PASSED- Measured by Digital Multimeter
PASSED- Measured by Observation
PASSED- Measured by Oscilloscope

## Slide 19
Communication Subsystem
Requirements
ESP32 must broadcast its SSID within 5 seconds of booting
HTTP server must serve HTML payload in under 500 ms
Web-submitted schedules must save to NVS without data corruption
SoftAP and web server must remain stable with two clients

## Slide 20
Description of project build and functional test results. You can choose to include a short (30s) video of your project here.

## Slide 21
Explanation of design and fabrication for all components o Introduction o Block Diagram  Requirements and verification with results

## Slide 22
Discussion of successes and challenges, as well as explanations of any failed verifications demonstrating and understanding of the engineering reason behind the failure

## Slide 23
Programming Issue

## Slide 24
Conclusions from the project: what did you learn, what would you do differently if you redesigned your project, etc.

## Slide 25
Timing & Response: Scent must trigger within 5s of the schedule, produce mist within 2s, and shut off within 1s if the lid opens.
Clock Sync: The system must sync with the client’s local time within 1 min of startup and maintain accuracy for all scheduled tasks.
Power & Safety: The supply must deliver 3.3V (300mA) and 5V (800mA) with <50mV ripple, including a 2.5V low-battery safety cutoff.
Conclusions

## Slide 26
Recommendations for further work.

## Slide 27
Optimize Design Flaws
Compact Integration
Future Work