Muscle Fatigue Interface

ECE445 Spring 2013

Design Review

TEAM 30: Shay Chen, Roman Levitas, Tushar Bhushan

TA: Lydia Majure

February 27th, 2013

Contents

I. Introduction..................................................................................................................................................3

1.1 Statement of Purpose….………………………………………………………………………..3

1.2 Introduction to EMGs…………………………………………………………………………..3

1.3 Objectives………………………………………………………………………………………..3

1.3.1 Goals…………………………………………………………………………………..3

1.3.2 Functions….…………………………………………………………………………..3

1.3.3 Benefits………………………………………………………………………………..3

II. Design..........................................................................................................................................................4

2.1 Block Diagrams............................................................................................................................4

2.2 Block Descriptions........................................................................................................................4

2.2.1 Overall Summary.........................................................................................................4		2.2.2 Data Acquisition Module.............................................................................................6		2.2.3 Preamplifier with Filtering..........................................................................................7 		2.2.4 Data Processing & Transmission Module..................................................................9		2.2.5 User Interface Module...............................................................................................12		2.2.6 Power Module.............................................................................................................13	2.3 Performance Requirements.......................................................................................................15

III. Requirements and Verification.............................................................................................................15

3.1 Requirements & Verification....................................................................................................15

3.2.1 EMG Module..............................................................................................................15 		3.1.2 LabView Module........................................................................................................15 		3.1.5 Storage Module...........................................................................................................16 		3.1.6 Display Module...........................................................................................................16

3.1.7 Power Module.............................................................................................................17

3.2 Testing Procedures…………………………………………………………………………….17

3.3 Tolerance analysis......................................................................................................................18

IV. Ethical Concerns.....................................................................................................................................18

V. Cost and Schedule.....................................................................................................................................19

5.1 Cost Analysis...............................................................................................................................19

5.2 Schedule.......................................................................................................................................20

5.3 Contingency Plan…………………………………………………………………...………….21

VI. Safety……………………………………………………………………………………………………21

VII. References..............................................................................................................................................23

I. Introduction

1.1 Statement of Purpose

This project was chosen because there presently are no devices that allow fitness enthusiasts to visually track their fatigue as they perform specific exercises. Such a device can aid in determining how effective their exercise is, as well as prevent overtraining. Knowing this information, one can program their rep/set scheme accordingly and thus increase the effectiveness of their workouts.  Moreover, there are many other biological applications including rehabilitation, physical therapy, and research. We have chosen to focus on the modularity of the device to make it inexpensive, transportable, and easy to maintain.

1.2 Introduction to EMGs

When muscles skeletal muscles contract, the number of motor units increases and the frequency of motor units firing increases4. When a motor unit, or motor neuron, fires, the nerve contracts the muscle fiber. An action potential is transmitted across the neuromuscular junction and into the muscle fibers. This causes an increase in electric potential which can be detected on the surface of the skin with surface electromyography (EMG). The activity of several motor units can be collected through the EMG. Depending on the size of the muscle, the electric potential from the muscle will be around 200-300µV.  Muscle fatigue is defined as a decrease in tension and power production caused by insufficient oxygen, depletion of energy, and the buildup of lactic acid. A decrease in pH of the muscle causes a decrease in the conduction velocity of the muscles which subsequently leads to a decrease in peak twitch tensions and an increase in contraction times4. This all leads to a decrease in median power frequency of the EMG signal, which our system uses as the indication of fatigue.

1.3 Objectives

1.3.1 Goals

●    	 Develop functional EMG with intuitive interface

●     Guarantee a quality commercial device

●     Allow optimization of different users

1.3.2 Functions

●     Instrumentational amplifier and filters to reduce artifacts in the EMG signals for accurate tracking

●     Analog to digital converter that also timestamps and stores the digitized information

●     Easy and intuitive interface

●     High intensity (6500 mcd) green LEDS lights to and beeper indicate fatigue level

1.3.3 Benefits

●     Inexpensive

●     Transportable

●     Prevents injuries from overtraining

●     Quantifiable measure of fatigue not available from human measurement

- Provides visual feedback & stimulus for physiotherapy patients to help them know when they are activating the muscle & chart their progress

1.3.4 Features

●     LEDs indicating level of fatigue

●     Sound indicating when muscle is near failure (maximum muscle hypertrophy)

●     Ability to vary the baseline median power frequency to accommodate different users/muscles

II. Design

2.1 Block Diagrams

2.2 Block Descriptions

2.2.1 Overall Summary

- Data Acquisition Module: This module obtains the electrical activity produced by skeletal muscles during exercise by using Electromyography (EMG) to detect changes in the electrical potential generated by muscle cells via surface electrodes.1 As muscular contractions are sustained, the spectrum of the electric signal shifts to the left (lower frequencies) as the firing rates of most motor units decreases, and the action potential of a fatiguing muscle attains a distinctly different shape than that of a fresh muscle (force twitches of motor units increase in amplitude during sustained contractions).2

  - Electrodes: The electrodes will serve as the transducer that will capture the muscle’s motor unit action potential (MUAP) and form our signal to be analyzed.

  - Attachment Armband: The armband will be used to hold the electrodes in place while the user is exercising.

- Preamplifier with Filtering Module: Since the peak amplitude of the signal picked up by the electrodes is usually in the range of 50μV - 1mV  3 (although 20-30 mV are not unusual, depending on muscle) we will need to amplify the signal to be able to analyze it. Furthermore, EMG signals are characteristic of having a bandwidth range from 0 to 500 Hz4 ; using a bandpass filter, we are able to eliminate noise from our amplified signal and have accurate measurements with no outliers.

  - Signal Amplifier: We will be using an op-amp to amplify the signal coming from electrodes.

  - Bandpass Filter: The amplified signal will go through another op-amp with a bandpass filter to filter out frequencies below 20Hz and above 500Hz.

- Data Processing & Transmission Module: This is the heart of the project. We will program a microcontroller that will analyze the analog signal coming in from the bandpass filter and perform the necessary calculations to be able to identify fatigue based on the median frequency of the EMG signal. Fatigue is described by a downward trend in median EMG signal. The module transmits the state of fatigue based on thresholds customized for the user around their baseline signals. The correct number of LEDs will light up corresponding to the state transmitted.

  - Microcontroller: We will be using a Labview compatible microcontroller that we can port Labview code to which will be able to correctly analyze the incoming signal.

  - Wireless Transmission Unit: This will communicate with a computer and transmit the analyzed data via bluetooth or wi-fi for logging and additional visualization.

- User Interface Module: This module consists of the elements that will notify the user of their muscle condition. A variety of colored LEDs will visualize the fatigue over time, a buzzer will indicate over-fatigue & a timer will be running for the duration of the workout being logged.

- Power Supply Module: Our power supply will consist of a 9 Volt source (a battery) that will be stepped down to the different voltages required by each module using DC/DC converters.

2.2.2 Data Acquisition Module:

Electrodes:

Inputs: Raw EMG pulses from muscle

Outputs: analog EMG signal

Description:

The two differential inputs are hooked up approximately an inch or two apart on the bicep. One is placed on the bulge on the bicep and another, an inch or two lower. Placing it on the bulge keeps away from tendons and other musculature. This is to establish the difference in voltage between the two points. The electrodes need to be close enough to be on electrically related tissue but not so close that there would be no difference in voltage. Because muscle fibers of motor units are distributed evenly, large muscle coverage is unnecessary. A reference electrode is placed on the shoulder and is connected to ground. As shown in Figure 1, it simply needs to be connected to electrically unrelated tissue. Before the testing, the user will wipe their tissue with alcohol to reduce excess noise and resistance from dirt and loose skin cells.

Figure 1: Representation of placement of electrodes on arm.4

2.2.3 Preamplifier with Filtering

AD622 Instrumentational Signal Amplifier:

Inputs: Raw EMG signal from electrodes

Outputs: Amplified raw EMG signal

Description:

The raw signal from the electrodes is first amplified by the AD622 instrumentation amplifier. This op-amp is able to measure a differential voltage on the skin and amplify it.It is quite effective in measuring the action potential of the motor neurons, reducing electromagnetic radiation, and allows for two electrodes. This chip is very easy to use and can be powered by a large range of power supply (+/-2.6V to +/-15V). It offers a good common mode rejection, a linear gain, and temperature stability.

The gain can be changed by varying the resistor Rg as seen in table 2.2.3.1. We amplify the signal raw signal by about 50V/V using a resistance of 1.02K. The op-amp will also amplify any residual noise around the circuit. Hence, we require an additional bandpass filter to remove this noise. In addition, there will be protection circuitry in the form of diodes and capacitors which ensure that the inputs to the amplifier cannot be greater than their turn on voltage. Capacitors would provide low impedance paths for voltage spikes and the resistors to dissipate the energy in the form of heat.

Rg = 50.5kW/(G-1)                                                      	(2.2.3.1)

Table 1:  REQUIRED VALUES OF GAIN RESISTORS

| Desired Gain | 1% Std Table Value of Rg, Ω | Calculated Gain |
| --- | --- | --- |
| 2 | 51.1 k | 1.988 |
| 5 | 12.7 k | 4.976 |
| 10 | 5.62 k | 9.986 |
| 20 | 2.67 k | 19.91 |
| 33 | 1.58 k | 32.96 |
| 40 | 1.3 k | 39.85 |
| 50 | 1.02 k | 50.50 |
| 65 | 787 | 65.17 |
| 100 | 511 | 99.83 |
| 200 | 255 | 199.0 |
| 500 | 102 | 496.1 |
| 1000 | 51.1 | 989.3 |

Bandpass Filter:

Inputs: Raw EMG amplified signal

Outputs: Filtered EMG amplified signal

Description:

The bandpass filter is built around the LM741 op-amp. It has overload protection built in and also has a good common mode rejection ratio. The bandpass filter will remove any outside noise and interference. Generally, 0-20 Hz is an unstable range indicating noise from the patient’s motion, the electrode placement, or the electrode cord. Frequencies above 500 Hz are generally noise from RF interferences.4 Thus, we require a bandpass filter ranging from about 20 to 500 Hz to reduce any noise that was amplified from the instrumentational amplifier. The bandpass filter will also have an additional gain of about 31V/V bringing the total gain of the preamp to 50*31= 1550V/V. Any more amplification tends to saturate the op amps. The resulting filtered EMG signal is then sent to the DAQ where it is then processed by the mock software/microprocessor.

Figure 2. External Circuitry of the EMG Module.

Low pass portion CL calculation:

High pass portion CH calculation:

Gain calculation

2.2.4 Data Processing & Transmission Module: NI Labview Programming Simulation

Figure 3. Labview code

Figure 3. shows the mock code written in Labview for signal processing. The data from the Data Acquisition module samples 2000 points at 2000Hz. The data is then sent through another digital bandpass filter with a low frequency cutoff of 20Hz and a high frequency cutoff of 500Hz. This second bandpass filter will act as a second order filter to ensure all unwanted frequencies are filtered out. The signal then goes through a bandstop filter which filters out frequencies 58-62Hz. Generally, 60Hz represents noise from the circuit and power lines.

Our system is essentially a third-order filter: one from the external circuitry, one from the Labview bandpass, and one from the Labview bandstop. After filtering the raw data, the square of the fourier transform of the data is taken. This represents the power spectrum - a range of frequencies and includes fast and slow motor units. The median frequencies is then taken and graphed. The median frequency is the indicator of muscle exertion. A typical median frequency of the bicep will start around 80Hz and failure tends to happen at 50Hz. Of course, each person is different and thus the system will account for it in an initial calibration stage.

As part of the mock interface, we have created a Fail Meter, which is the software version of the fatigue indicator based on the median frequency. Based on experimental data, a range of the Fail Meter is created between 50Hz and 100Hz. Multiple ranges will be created in increments of 10Hz to feed into the Fail Meter which will update every at data point. The Fail Meter is a software representation of the LED’s that will be implemented. Also, when the Fail Meter is almost filled, the beeper will sound.

Figure 4. Full simulation of preamp + Labview code.

Microcontroller:

Inputs: Amplified signal from the Pre-amp

Outputs: Signals to LEDs, Signal to Sound Buzzer

Description:

The microcontroller will be used to process the signal coming from the filter part of the circuit. This will be connected to three additional components, that is, the LED display, timer, and the sound buzzer. The microcontroller will also calibrate the device based on the user, since each user has their own baseline signals.

For the microcontroller, we will be using the Arduino Uno R3 board. This is a relatively portable board and performs all the signal processing we will need to determine the user fatigue. It will take in the output of the hardware based filter as input and it outputs a signal to the LED’s or the sound buzzer as necessary.

The microcontroller itself will take the input signal and will run a Fast Fourier Transform (FFT) on the given input. This will be implemented using the FFT library found on a tutorial website for the Arduino board. 6 The output of the FFT will be analyzed to get the frequency of the input signal as well as the amplitude. These values can then be used to drive the LED’s depending on the range of values of amplitude and frequency as required. Additionally, the onboard XBee wireless interface can be used to wirelessly transmit data about the amplitude peaks onto a computer. This information can then be graphed to give a visual representation of a user’s muscle fatigue.

The Arduino Uno R3 package is shown below:

Figure 5: Arduino Uno R3

We will be using the pins for analog input and digital output to and from the microcontroller with the following pins assignments:

Figure 6: Arduino Pin Schematic

Wireless Transmission Unit:
Inputs: Signal from Arduino Board
Outputs: Wireless transmitter
Description:
The onboard XBee wireless interface can be used to wirelessly transmit data about the amplitude peaks onto a computer. This information can then be graphed to give a visual representation of a user’s muscle fatigue.

2.2.5 User Interface Module:

LED Display:

Inputs: Signals to LED

Outputs: Visible feedback

Description:

The LED display will take its inputs from the microcontroller, and will give the user a visual representation of the amount of stress/fatigue that is being placed on the muscles.

Sound Buzzer:

Inputs: Signal to Sound Buzzer

Outputs: Aural feedback

Description:

Like the LED display, the sound buzzer too, takes its inputs from the microcontroller. The purpose of the buzzer is to give the user an additional source of stimuli in order to warn the user when their muscles are near failure. These parameters vary from person to person; however, the microcontroller will be calibrated per user.

Figure 7: PIN Diagram for the Sound Buzzer

2.2.6 Power Module:

Inputs: 9 volt battery

Outputs: power connections at specified voltages

Description:

The power supply module is at the heart of powering our three main components: the microcontroller and the two amplifiers. The microcontroller itself will power the User Interface module using its output logic signals dropped over resistors. Our microcontroller, the Arduino, also requires the largest voltage in our circuit: 9 Volts. The amplifiers operate well with a standard 5 volt supply, like most TTL chips, and for that we will step down the 9 volt supply using a 78SR-5/2-C: DC/DC Converter as shown in the block diagram. The table below summarizes the voltage allowances of each part and what we will be implementing.

\

Figure 6. Power supply

Table 2: Voltage Requirements

| Device | Voltage(s) Tolerances | Voltage Used |
| --- | --- | --- |
| Arduino UNO R3 | 7-12V | 9V |
| LM741 Op-Amp | ±22V | ±5V |
| AD622 Instrumentation Amplifier | ±2.6V - ±18V | ±5V |

Based on the Voltages we decided to use and the devices’ maximum current draw, we expected a maximum energy drawn of 132.5 mW:

Power Consumption:
LM741 (Max) Power: 50mW
AD622 (Max): 1.5mA X 5V = 7.5mW
Arduino (Max): 25mA X 5V = 75mW
				= 132.5mW

If we use an Energizer 522 9 volt battery we can expect a life time of between 400 and 500 hours as indicated by the data sheet:

Figure 7. Energizer 9 V lifetime

2.3 Performance Requirements

- The completed product should be able to withstand human body temperatures while exercising and should not affect the functionality or stability of the components.

- The enclosure and design should be small enough for convenient usage and transport, including the power supply and data communication.

- The device should be able to calibrate itself to different users & their baselines to provide pertinent data.

- LEDs are correctly correlated to muscle fatigue and the sound buzzer is activated once at a predefined threshold.

III. Requirements and Verification

| Block | Requirements | Verifications |
| --- | --- | --- |
| Electrodes | a. Electrodes can detect a voltage differential. b. Electrodes read voltage values of 50 µV - 50mV on bicep. | a.1 Apply a 10mV PtP sinusoidal voltage from function generator across two electrodes. a.2 Hook output to oscilloscope, Oscilloscope should read 10mV PtP b.1 Wipe skin with alcohol swab. Apply eletrogel to increase coupling. b.2 Connect the electrodes correctly onto bicep skin b.3 Connect the output of the electrodes to an oscilloscope b.4 Measure the values to ensure a voltage range of 50 µV - 50mV. |
| Instrumentional Amplifier | a. The amplifier needs to have a gain of 30 V/V +/-1.5V/V. | a.1 Check op-amp is powered a.2 Input from the function generator a sine wave of amplitude 10mV. Check the output amplitude for the required gain. Resulting sine amplitude should be 500mV +/-%5 a.3 If the gain is incorrect, use an ohmmeter to test the resistor. |
| Bandpass Filter: | Amplifier passes signals. b. The noise needs to be <2.5µV RMS c. Frequency response needs to be: Low Pass(-3 dB) : 10-20Hz High Pass(-6 dB) : 480-500Hz d. Additional gain of 30V/V e. Bandstop filters out 58Hz-62Hz | a.1 Check op-amp is powered a.2 Use a frequency generator to input white noise. Use the output of the filter to view the Fourier transform on the oscilloscope b.1 Input white noise into the filter (0Hz-1kHz) and view the Fourier transform of the output on oscilloscope b.2 Use oscilloscope to check noise c. Input white noise into the filter (0Hz-1kHz) and view the Fourier transform of the output on oscilloscope d.1 Check op-amp is powered d.2 Input from the function generator a sine wave of amplitude 10mV. Check the output amplitude for the required gain. Resulting sine amplitude should be 300mV +/-%5 d.3 Use ohmmeter to measure resistors. Tolerance +/-5% e1. Input white noise into the filter (0Hz-1kHz) and view the Fourier transform of the output on oscilloscope e2. Frequency response needs to be stopband (-3dB) 58Hz-62Hz |
| Microcontroller | a. Microcontroller needs to be able to take the signal outputs of the bandpass filter as inputs b. It should be able to perform signal processing on the input waveforms c. Based on these values it should drive the LED display or the sound buzzer as required. d. Operating voltage range of 2.4-5.5V e. DAQ has a resolution of 10bits | a.1 Take the output of the bandpass filter and connect it to the microcontroller. a.2 Make sure DAQ is properly connected and powered. a.2 View the output on LabView to ensure correct connection. b.1 The output of the processed signal should be free from noise, i.e. there minimal frequencies - less than 6db of >500Hz and <10Hz b.2 This can be verified using LabView to analyze the output signal c.1 The output signal of the microcontroller drives power to the LEDs c.2 Once the median frequency is the baseline. d.1 Connect to power supply to the power supply at 3V, and send a test signal from the function generator and compare with the output signal from the oscilloscope. Voltage tolerance +/-5% e.1 Feed in a 10mv PtP sine wave e.2 Resulting digital signal should have a resolution of 10bits |
| LED Display | a. LED Display should interact correctly with user | a.1 Connect LEDs to Vcc and Ground via resistor a.2 Run 2V source from a function generator and see if they light up. Voltage tolerance +/-5% |
| Sound Buzzer: | a. Buzzer beeps when failure occurs b. Buzzer operates at a voltage range of 2V | a.1 Ensure buzzer is powered a.2 Ensure buzzer is connected to microcontroller a.3 Microcontroller sends signal when failure occurs b.1 Run 2V source from a function generator and see if the buzzer beeps. Voltage tolerance +/-5% |
| Data logger | a. Implemented via software, run several trials b. Bit error threshold of .1% | a.1 Check if the micro-controller is sending signals to the wireless output. a.2 Check if software is receiving data. a.3 Check software code for errors. b.1 Output a string of 1000 bits. b.2 Resulting data logged on PC should have less than 10 errors. |
| Power Module | a. Supply a constant 9 volt signal to power the microcontroller b. The supply is stepped down correctly to 5 Volts | a. Use a multi-meter to ensure 9 Volt battery supply and use an oscilloscope to check if the voltage is steady over time. b. A multimeter will be used to check the output of the DC/DC convertor and that the received voltage at the amplifiers is 5 volts ± 0.20V |

3.2 Testing Procedures

The initial testing that we will be doing will be on the EMG amplifier. This test will ensure that we get correct performance from the circuit. We will do this testing on an oscilloscope and test the EMG on the bicep muscles with different degrees of strain. This will establish the necessary settings for the EMG test circuit. We can then confirm what potential the bicep performs at and compare that to the expected EMG readings that others have found. We will have a table displaying muscle potential versus strain so that we can see what weight should be used for a particular person. Time permitting, we will try to connect the EMG to other muscle parts.

3.3 Tolerance Analysis

Tolerance analysis will be performed mainly on the preamp itself. The EMG is the primary component to our device and without an accurate signal from the preamp, everything becomes moot. Our EMG signal must be properly amplified, and properly filtered. We plan to use an AD622 Instrumentation Amplifier which has a gain relationship or Rg=50.5Kohm/(G-1). We will vary Rg and test its tolerance extremes. Also, we will test the resistors in the bandpass filter and note the change in gain. Overall, we want to find an adequate raw EMG signal for further digital signal processing.

IV. Ethical Concerns

Our project aims to provide users with the ability to measure their muscle fatigue while undergoing strenuous activities. While this is aimed towards the gym going user, it also has other uses in the field of medicine, to prevent patients with muscle damage from straining themselves excessively.

The failure of any electrical instrumentation making direct or indirect galvanic contact with the skin can cause a potentially harmful fault current to pass through the skin of the subject. This concern is less relevant in devices that are powered exclusively by low voltage (3-15 V) batteries. To ensure safety, the subject should be electrically isolated from any electrical connection (to the power line or ground) associated with the power source. This isolation is generally achieved in one of two ways: either through the use of optical isolators or through the use of isolation transformers. Both approaches are satisfactory, but both require careful consideration for not distorting the EMG signal. This is especially true when a transformer is used. This isolation provides the added benefit of reducing the amount of radiated power line noise at the electrode detection surfaces.

These functions increase the safety of the user, which is consistent with the first code of the IEEE code of ethics.

1. to accept responsibility in making decisions consistent with the safety, health, and welfare of the public, and to disclose promptly factors that might endanger the public or the environment;

One of the factors necessary to avoid muscle fatigue is calculating an accurate baseline frequency for each individual user, and based on these values, the muscle fatigue can be gauged. Throughout the development of this device we will follow the third code closely, and make estimates about the baseline from real data.

3. to be honest and realistic in stating claims or estimates based on available data;

While working on the project, we will build an environment that promote engineer

professionalism, which welcomes constructive and honest criticisms, acknowledges errors, assists

peer workers with their professional and academic developments, and credits appropriate

contributions, as cited in the 7th and 10th codes of the IEEE Code of Ethics:

7. to seek, accept, and offer honest criticism of technical work, to acknowledge and correct

errors, and to credit properly the contributions of others;

10. to assist colleagues and co-workers in their professional development and to support them in following this code of ethics.

Additionally, we will ensure that the readings given out by the interface are as accurate as possible and will not provide them with false information, in order to avoid any damage that may occur to a user’s health, as per the ninth code of ethics.

9. to avoid injuring others, their property, reputation, or employment by false or malicious

action;

V. COST AND SCHEDULE

5.1.0 Labor

We estimate working hour ranges from 15 to 25 hours per week with a total duration of 10 weeks to complete this project.

| Name | Hourly Rate | Total Hours Invested | Total = Hourly Rate x 2.5 x Total Hours Invested |
| --- | --- | --- | --- |
| Shay Chen | $25 | 200 | $12,500 |
| Roman Levitas | $25 | 200 | $12,500 |
| Tushar Bhushan | $25 | 200 | $12,500 |
| Total |  | 600 | $37,500 |

5.1.1 Parts

| Part | Cost Per Unit | Quantity | Total Cost | Order Status |
| --- | --- | --- | --- | --- |
| Skin Electrodes | $.50 | 4 | $2 | University |
| Surface Mount PCB | $30 | 1 | $30 | University |
| Wires | $2.50 | 2 spools | $5 | University |
| Resistors & Caps & Diodes | $.10 | 30 | $3 | University |
| Batteries | $.30 | 8 | $2.40 | Ordered |
| Battery holder | $3.17 | 1 | $3.17 | Not Ordered |
| LED’s | $.15 | 10 | $1.50 | Not Ordered |
| Bar Graph Disp. | $4.97 | 1 | $4.97 | Not Ordered |
| Metal Enclosure | $40 | 1 | $40 | Not Ordered |
| Beeper | $5.30 | 1 | $5.30 | Not Ordered |
| A/D Converter | $20 | 1 | $20 | University |
| AD622 op-amp | $2.56 | 1 | $2.56 | University |
| LM741 op amp | Free Sample | 1 | - | Ordered |
| Arduino UNO R3 | $45.65 | 1 | $45.65 | Ordered |
| 78SR-5/2-C DC/DC Converter | $12.36 | 1 | $12.36 | Not Ordered |
| Misc. | $20 | 1 | $20 | Not Ordered |
| Total |  |  | $197.91 |  |

5.1.2 Grand Total

$37,500 in labor + $197.91 in parts = $37,697.91

5.2 Schedule

| Week | Task | Responsibility |
| --- | --- | --- |
| 2/3 | Finalize and hand in proposal | Shay |
|  | Research biopotential amplifiers | Roman |
|  | Research ideas for implementation | Tushar |
| 2/10 | Design biopotential amplifier | Shay |
|  | Design bandpass and anti-noise filters | Tushar |
|  | Mock Design Review | All |
| 2/17 | Simulate design and start drawing out board design | Shay |
|  | Design rep/set tracking software | Roman |
|  | Order parts for EMG | Tushar |
| 2/24 | Design Review | Roman |
|  | Assemble power supply and EMG for testing | Shay |
|  | Lay out PCB board for display module and beeper | Tushar |
| 3/3 | Order display module and PCB’s | Shay |
|  | Optimize EMG for multiple users/muscles | Roman |
|  | Testing/Debugging EMG | Tushar |
| 3/10 | Assemble PCBs for display | Tushar |
|  | Testing/Dubugging EMG | Shay |
|  | Assemble PCBs with tracking software | Roman |
| 3/17 | Prepare mock-up demo | Roman |
|  | Testing/Debugging EMG | Shay |
|  | Finalize PCB | Tushar |
| 3/24 | Mock up demo | Roman |
| 3/31 | Mock up presentation | Roman |
| 4/7 | Ensure Completion. Last minute debugging & finalization | Shay |
| 4/14 | Prepare Demo and Presentation | Tushar |
| 4/21 | Demo | Shay |
| 4/28 | Presentation | Roman |
|  | Final paper | Tushar |
|  | Checkout | Roman |

5.3 Contingency Plan
· In the event that we are unable to use a microcontroller to perform the signal processing that we need, we will implement the functions using LabView.
· If the data logging software is unable to be implemented using the wireless transmission unit, we plan on using a wired interface through LabView.

VI. Safety

Our project will have voltages up to 9 V and 25-30 mA of current which could lead to some unpleasant results. Table 3 summarizes the effects of electrical currents on the human body according to. Since or project deals with currents of in the range of 9-62mA, we must ensure we take all the necessary precautions when assembling our design.

Table 3: Electricity’s Effect on Human Body

| Effects | Current, mA |  |
| --- | --- | --- |
|  | Direct Current | Alternating Current |
| Sd | DC | 60 Hz |
| Slight sensation on hand | 1 | 0.4 |
| Perception threshold, median | 5.2 | 1.1 |
| Shock—not painful and no loss of muscular control | 9 | 1.8 |
| Painful shock—muscular control lost by 1/2% | 62 | 9 |
| Painful shock—let-go threshold, median | 76 | 16 |
| Painful and severe shock—breathing difficult, muscular control lost by 99.5% | 90 | 23 |

If something goes wrong while we’re working, we could end up experiencing from no more than a shock to some loss of muscular control.

We will primarily be working in the Electric Machinery Lab (Everitt Lab 50) which has some of its own rules and warnings, including:

- GROUND! Be aware of which connections are grounded, and which are not.

- RATINGS! Before applying power, check that the voltage, current, and power levels you expect to see do not violate any device ratings.

- HEAT! Small parts can become hot enough to cause burns with as little as one watt applied to them.

- CAREFUL WORKMANSHIP! Check and recheck all connections before applying power.

Plan ahead: consider the effects of a circuit change before trying it. Use the right wires and

connectors for the job, and keep your bench neat.

- WHEN IN DOUBT, SHUT IT OFF! Do not manipulate circuits or make changes with

power applied.

- LIVE PARTS! Most semiconductor devices have an electrical connection to the case. Assume that anything touching the case is part of the circuit and is connected. Avoid tools and other metallic objects around live circuits.

Keep beverage containers away from your bench.

The laboratory is equipped with an emergency electrical shutoff system that can be activated by pushing a button located on the benches. The benches and equipment are also fitted with circuit breakers that will stop fault current when necessary and prevent damage.

.

VII. References

1) Electromyography at the US National Library of Medicine Medical Subject Headings (MeSH)

2) Moritani, T., & Muro, M. (1987). Motor unit activity and surface electromyogram power spectrum during increasing force of contraction. European Journal of Applied Physiology, 56, 260-265.

3) Nakra, B. C., and K. K. Chaudhry. "Measurements of Bioelectric Potentials." Instrumentation Measurement and Analysis. 3rd ed. New Delhi: Tata McGraw-Hill, 2009. p. 418.

4)  De Luca, Carlos J. “Surface Electromyography: Detection and Recording.” Delsys, 2002.

http://www.delsys.com/Attachments_pdf/WP_SEMGintro.pdf

5) Bigland-Ritchie B (1981) EMG/force relations and fatigue of human voluntary contractions. Exerc Sport Sci Rev 9:75-- 117

6) Dieder. “Fast Fourier Transform (FFT) Tutorial”. 05 October 2010. Web.

http://www.arduinoos.com/2010/10/fast-fourier-transform-fft/