## Slide 1: Intelligent Home Security System
Group 13
Danyu Sun
Ge Shao
Jiaxuan Zhang
Yuxin Xie
ECE 445 Senior Design
May 25, 2025

## Slide 2: Introduction: Problem Overview
Growing elderly population living alone
  → Increased safety concerns at home
Fall accidents are common and dangerous
  → Can cause serious injury or death if not detected quickly
Lack of real-time, responsive monitoring in home settings
To reduce the risk and anxiety associated with unattended falls

## Slide 3: Introduction: Solution
Main Features:
  YOLO-based human and posture recognition
  Real-time tracking with distance sensors
  Immediate audible and visual alarm after a fall
Benefit:
  Improves home safety and supports remote elderly care scenarios.

## Slide 4: Objective
Human detection
Autonomous trace
Fall detection
Automatic alert

## Slide 5: System Overview
Notes: block diagram

## Slide 6: System Overview
Notes: schematic circuit diagram

## Slide 7: Design of Robot Base
Chassis
DC Gear Motor (×4)
Rubber Wheel (×4)

## Slide 8: Design of Robot Base
Arduino UNO R3
PWR (Power) Board
Battery Pack

## Slide 9: Design of Human Tracking
Ultrasonic Sensor
HC-SR04

## Slide 10: Design of Human Tracking
Infrared Sensor
E18-D80NK (x2)

## Slide 11: Design of Video Capture
Gimbal Frame
Servo Motor (x2)
Wi-Fi Module with Wireless Antennas

## Slide 12: Design of Video Capture
Camera
Video Stream URL | http://192.168.1.1:8080/?action=stream
Video Encoding | MJPEG
Resolution (Pixels) | 640 × 480
Frame Rate | 15 FPS
Field of View | 30° (horizontal)
Power Supply | 5V USB

## Slide 13: Design of Alert
LED module
Active Buzzer

## Slide 15: Fall Detection Algorithm - Choose
YOLO-NAS is the BEST!

## Slide 16: Fall Detection Algorithm - Core Logic of YOLO-NAS
Search
  Automatically explores and combines new architectures from a pool of building blocks.
Evaluate
  Cyclically evaluates each architecture for accuracy, quantization stability, and real hardware speed.
Select
  Outputs the best model, which uses reparameterization to fuse its blocks for the fastest possible inference speed.
Notes: 左侧 (Search): 代表“神经架构搜索”（NAS）的起点，算法从大量的可选构建块（Search Space）中进行探索。
中间 (Evaluate): 这是搜索循环的核心，对每一个候选架构都进行三个关键维度的评估：准确率 (Accuracy)、量化稳健性 (Quantization Robustness) 和 硬件延迟 (Hardware Latency)。评估结果会反馈给控制器，以指导下一步的搜索，使其越来越好。
右侧 (Select): 经过充分的搜索和评估，最终输出的最佳模型就是YOLO-NAS。这个模型最大的特点就是我们提到的“再参数化”（Reparameterization）：在训练时使用更复杂的结构来提升性能，在部署推理时则会融合成极其简单和高效的结构。

## Slide 17: Fall Detection Algorithm - Finetune on Our Dataset
To rapidly adapt YOLO-NAS for high performance detection on our own dataset.
  Frozen Layers: Froze all early layers to retain their powerful, general feature-detection capabilities.
  Targeted Training: Trained only the final two layers using our small, custom dataset.
This strategy successfully produced a highly accurate and specialized detector with minimal training time, proving its effectiveness even on a small dataset.

## Slide 18: Requirement & Verification
Requirement 1: Low-latency video transmission
Verification: A subject performed predefined motions while the video stream was logged. Timestamps were compared between real-world action and display output. 30 trials were conducted.
Result: Average latency was 148 ms, maximum 172 ms, both within the <200 ms threshold.
Notes: It proved that the video transmission of our video capture design has a reasonable delay which can be ignored.

## Slide 19: Requirement & Verification
Requirement 2: Accurate human movement detection
Verification: A person walked in front of the robot at 0.5 m, 1 m, and 1.5 m, at approximate speeds of 0.5 m/s, 1.0 m/s, and 1.5 m/s respectively. For each condition, 10 tests recorded whether the sensors triggered correct tracking behavior.

## Slide 20: Requirement & Verification
Result:
IR detection is reliable within 1 m across all tested speeds.

## Slide 21: Requirement & Verification
Requirement 3: Real-time video input processing
Verification: Live video was streamed to the system while a subject performed timed fall gestures. Timestamps were recorded at the moment of physical action and upon system recognition. 30 trials were conducted under consistent indoor lighting.
Result: The average processing latency was 1.5 s, with all results below 5 s, satisfying the <8 s real-time requirement.

## Slide 22: Requirement & Verification
Requirement 4: Human recognition accuracy over 90% under indoor lighting
Verification: Subjects wearing common clothing colors (black, white, gray, blue, red) were tested against three background colors (white wall, brown wood, green curtain). Each combination was tested 5 times while subjects walked or stood. A total of 75 trials were manually labeled and compared with system outputs.

## Slide 23: Requirement & Verification
Result :
Total tests: 75
Overall average accuracy: 97.6%
Minor accuracy drops occurred when white clothing blended with light-colored backgrounds, but overall recognition performance exceeds the requirement threshold.

## Slide 24: Requirement & Verification
Requirement 5：fall detection accuracy over 80% under indoor lighting
Verification: Subjects wearing common clothing colors (black, white, red) were tested against two background colors (white wall, brown wood). The subjects respectively used the forward fall, backward fall and lateral fall (facing and back to camera) methods, each for 5 times. A total of 120 trials were manually labeled and compared with system outputs.

## Slide 25: Requirement & Verification
Total tests:
Overall average accuracy: 85.0%
Minor accuracy drops occurred when white clothing blended with light-colored backgrounds, or when subjects turn the back to the camera. But overall recognition performance exceeds the requirement threshold.

## Slide 26: Requirement & Verification
Requirement 6: Quick response of the alert subsystem
Verification:Fall events were simulated, and the activation time of both the LED and buzzer as well as the duration of the buzzer alarm, were measured using a stopwatch. Each test was repeated 10 times under consistent indoor conditions.

## Slide 27: Requirement & Verification
Result: The LED activated in an average of 0.52 seconds (range: 0.41–0.62 s), and the speaker in 0.51 seconds (range: 0.47–0.55 s), with the single alarm lasting an average of 10.7 seconds (range: 10.3–11.1 s).
Both components met the system requirements: activation within 1 second and speaker duration exceeding 10 seconds.

## Slide 28: Challenges & Solution
Function | Challenge | Solution
Video Capture | The camera's field of view is limited, preventing it from capturing the entire human body at some specified distances. | Select a camera with a wider field of view. Or increase the number of cameras and reposition them to ensure full-body coverage.
Human Tracking | The sensors have insufficient maximum detection ranges for human tracking and show a great disparity (e.g., IR sensor: 80 cm, ultrasonic sensor: 400 cm). | Replace the sensors with those having larger detection ranges and matching each other.
Human Detection ( for tracking ) | The sensors lack the capability to distinguish between human subjects and other objects. | Use human detection camera or algorithm to accurately differentiate humans from other subjects
Center Control Unit | Fall detection algorithms on the PC, not directly stored on the robot itself | Replace the Arduino Uno and the external PC with an onboard computing unit like Raspberry Pi, which can run fall detection algorithms locally.

## Slide 29: Failure & Solution
Failure: We observed that sometimes for a period of time after the car started but before the fall detection function was activated, the LED lights and the buzzer would react abnormally.
Explanation: When Arduino powers on, it may read garbage data or startup noise from the serial port, causing the LED and buzzer to trigger incorrectly.
Solution：Clear the serial buffer in setup() and only accept valid characters '1' or '0' when reading serial input

## Slide 30: Other tests -Brownout Recovery Test
Using a variable DC power supply, we lowered the system’s input voltage gradually from 7.4V to 6.0V.
At each voltage level, we observed:
Whether the Arduino booted correctly;
Whether the motor and WiFi module functioned as expected;
Whether the LED and buzzer could respond to fall detection signals;
Whether serial communication remained stable
Notes: To evaluate whether the system remains functional when the input voltage drops below the recommended minimum (e.g., 6V), simulating battery drain or degraded supply conditions.

## Slide 31: Other tests -Brownout Recovery Test
Result：
7.4V ~ 6.5V: system operated normally.
6.2V: motors began showing signs of reduced torque, but system remained functional.
6.0V: WiFi module failed to maintain connection intermittently, and voltage dropped below 4.7V at the Arduino VCC pin.
< 6.0V: system became unstable; serial communication often failed.
Improvement：
Develop low battery alert signal to remind users to recharge and avoid system instability under low battery conditions

## Slide 32: Future Hardware Development
Improve tracking accuracy: Increase the number of sensors and replace distance IR sensors with thermal IR sensors.
Enhance monitoring coverage and image quality: Upgrade to multiple cameras or use a camera with a wider FOV and higher frame rate.
Achieve fully autonomous operation: Add an automatic charging system for self-docking and battery recharge.
Ensure stable tracking in different lighting conditions: Use sensors that are more resistant to ambient light interference.

## Slide 33: Future Software Development
Remote Alarm Notification: Integrate wireless communication to send fall alerts to family members in real time, enabling remote assistance and timely intervention.
Extended Behavior Recognition: Expand the system to detect other risky behaviors such as prolonged inactivity, unusual sitting/lying postures, or wandering, further enhancing home safety monitoring.
Improved User Interface & Interaction: Develop a user-friendly mobile app for easier control and feedback, especially tailored for elderly users and family members.

## Slide 34: Demo Video
Tracking demo

## Slide 35: Demo Video
Falling demo

## Slide 36: Thank You for Listening! Q & A