## Slide 1: A Compact Material Modulus Measurement Instrument​ ECE445 Senior Design Final Presentation
Group1:
Kongning Lai, Yunzhi Lu, Ziyi Lin, Tianyu Fu
TA: Zheyi Hang
Supervisor: Prof. Huan Hu
2025.5.25

## Slide 2
CONTENT
1
2
3
4
Introduction
Background and Objective
Block Diagram
Design and Verification
Success and Challenges
Development and Business
Mechanical Module
Analog Electrical Module
Control Module
Software Module
Future Work
SWOT Analysis
Notes: 停留3秒，让观众看一下

## Slide 3
Background and Objective
Background
Soft materials widely used in biomedicine, soft robotics, wearable devices
Mechanical characterization is essential
Existing systems struggle to balance precision, cost, and accessibility
Objective
A macro-scale AFM-inspired instrument
Build user-friendly software for real-time measurement and analysis
An affordable, easy-to-use tool for education and screening
Image source: Liu et al., Trends in Chemistry, 2020 [1].

## Slide 4
Block Diagram
Notes: Mechanical Module: A stepper motor drives a lead screw for precise movement. The cantilever applies force to the sample.
Analog Sensing: A strain gauge in a Wheatstone bridge detects deformation. The signal is converted and amplified by HX 711 Adc
Control: An STM32 microcontroller handles motor control, data collection, and communication. A user-friendly interface controls motion and displays real-time force-displacement data.

## Slide 5
Mechanical Module
Requirements | 
Step Resolution | <1.0 μm per step with 2 mm lead screw
Spring Constant and Sensitivity | k = 10 - 15 times higher than sample Vertical Displacement > 0.1 mm
Table 1. Different Cantilever Design
(a)
(b)

## Slide 6
Mechanical Module
Verification and Results | 
Step Resolution | 0.3 μm/ Step
Spring Constant and Sensitivity | = 267.7 N/m >10 Deflection under 0.1 N load is 0.31 mm
Verification and Results | 
Step Resolution | 0.3 μm/ Step
Spring Constant and Sensitivity | 
Table 2. Mechanical Properties of Different Cantilever Design
(b)
(c)
(a)

## Slide 7: Analog Electrical Module
cantilever
F
Notes: The Wheatstone bridge (Figure 4 left) is used to measure the unknown resistance of the strain gauge. It is constructed using three 220 Ω resis tors, forming a balanced circuit when the gauge is unstressed. Any bending of the gauge leads to a change in its resistance Rx, resulting in a measurable voltage Vg that reflects strain. To enhance the signal-to-noise ratio, the output voltage is amplified and the then passed to an ADC for digital processing and storage.

## Slide 8: Analog Electrical Module
Selection of the Amplifier
V
S
TI INA125
HX 711
Key Reasons for Choosing
HX711 + STM32 over INA125 + STM32​
Higher Precision at Lower Cost​
Susceptibility to EMI/RFI​
PCB overview
Notes: After measuring the voltage output of the wheatstone bridge the output without the amplification was only around 1-2mv, we need an amp. We tried both and dicided to use hx711
Next part we will talk about serial communication

## Slide 9: Control Module
Table 3. STM32 Pin Assignment Summary
(a)
(b)
Notes: The Wheatstone bridge (Figure 4 left) is used to measure the unknown resistance of the strain gauge. It is constructed using three 220 Ω resis tors, forming a balanced circuit when the gauge is unstressed. Any bending of the gauge leads to a change in its resistance Rx, resulting in a measurable voltage Vg that reflects strain. To enhance the signal-to-noise ratio, the output voltage is amplified by an INA125 amplifier (Figure 4 right), and the amplified signal is then passed to an ADC for digital processing and storage.

## Slide 10: Control Module
Motor Pulse (PWM via Timer) | HX711 Clock (Manual GPIO Toggle)
Hardware-driven, efficient | Software-driven, flexible
Precise frequency & duty cycle | Precise timing control needed in code
No CPU overhead after setup | Requires tight CPU control
Ideal for continuous waveform | Ideal for low-speed serial interface
Comparison: Timer-based PWM vs GPIO Manual Clocking
(a)
(b)
Notes: The Wheatstone bridge (Figure 4 left) is used to measure the unknown resistance of the strain gauge. It is constructed using three 220 Ω resis tors, forming a balanced circuit when the gauge is unstressed. Any bending of the gauge leads to a change in its resistance Rx, resulting in a measurable voltage Vg that reflects strain. To enhance the signal-to-noise ratio, the output voltage is amplified by an INA125 amplifier (Figure 4 right), and the amplified signal is then passed to an ADC for digital processing and storage.

## Slide 11: Control Module
HX711 Noise Analysis:
Noise Level, Std Dev: 708.2 μN
Experimental force range: > 100,000 μN
Notes: The Wheatstone bridge (Figure 4 left) is used to measure the unknown resistance of the strain gauge. It is constructed using three 220 Ω resis tors, forming a balanced circuit when the gauge is unstressed. Any bending of the gauge leads to a change in its resistance Rx, resulting in a measurable voltage Vg that reflects strain. To enhance the signal-to-noise ratio, the output voltage is amplified by an INA125 amplifier (Figure 4 right), and the amplified signal is then passed to an ADC for digital processing and storage.

## Slide 12: Software Module

## Slide 13
Motor control
  RUN/STOP button
  DIR button
Data collection
  F-Z conversion:
Add mass values to the strain gauge ranging from 0g to 50g in 5g intervals
Combine all corrected sensor readings
Use Linear Regression to fit a model
  Real-time display
  Save
Serial Port Communication

## Slide 14: Data Processing
Pipeline
E_eff = 1.6062 MPa
R² = 0.9857
R = 10.00 μm
Saved F–Z Input
Contact Point Estimation (via CNN)
Contact Model Fitting
Elastic Modulus Calculation
Contact point estimation
  Manually labeled contact points (displacement, force) used as ground truth
  Loss function: MSE between predicted and labeled contact points

## Slide 15: Sucesses & Challenges

## Slide 16: Sucess
ERROR < 10% !

## Slide 17: Challenges
Accuracy
Displacement measurement accuracy can be improved
Limited Generalization
Neural network trained on a small dataset
Material Property
Measured elastic modulus is affected by tip speed due to soft material viscoelasticity [2]
Contact Model
Hertz model assumes ideal contact; real conditions are more complex

## Slide 18: Future Work
Infrared sensor for displacement feedback
01
02
03
One-button Test Execution with Configurable Parameters
Larger ML dataset for model robustness

## Slide 19
SWOT Analysis
Resolution:
Cannot compete with nanometer-level resolution;
may face challenge in high-end lab settings.
Limited Sample Validation:
Primarily tested on PDMS; foam materials lacked ground truth modulus for comparison.
Cost-Effective Screening
Can be marketed for rapid screening in materials research labs, especially in budget-constrained environments.
Low Cost and Accessibility:
below 1000 RMB for hardware
(excluding labor)
affordable compared to
commercial systems.
S
W
O
T

## Slide 20
Reference
[1] J. Liu, Y. Gao, Y.-J. Lee, and S. Yang, "Responsive and foldable soft materials," Trends in Chemistry, vol. 2, no. 2, pp. 107–122, 2020.
[2] J. H. Owen, "Toward a better modulus at shallow indentations-Enhanced tip and sample characterization for quantitative atomic force microscopy," Microsc. Res. Tech., vol. 86, no. 1, pp. 84–96, Jan. 2023, doi: 10.1002/jemt.24261.

## Slide 21: Thank You! Questions Welcome!